## GP8600-PWM-CN
DAC模组精度计算
GP8600模块 | [5 empty] | 拟合校准算法 y=c*x+o 理论程序参数x,变换后的程序参数y 变换系数c，变换偏移o值 |  |  |  | 计算步骤 1、计算斜率 k: =SLOPE(D11:D15,B11:B15) 2、截距 b: =INTERCEPT(D11:D15,B11:B15) （步骤1、和2舍去了第一个和最后一个端点数据，因为当端点数据误差过大时会影响其他数据的拟合效果） 3、计算变换系数 c: =(理论最大电压/最大程序参数)/k 4、计算变换偏移o值 o: =-b/k 5、计算变换后的程序参数y=c*x+o
参考datesheet参数：±0.2%满量程FSR，电流/电压线性度误差0.05%typ
校准说明：校准需要搭配MCU，所以这里的校准数据及误差只做参考，DAC模块需要客户自行进行校准。
PWM输入 | 实际斜率k | 实际截距b | 变换系数c | 变换偏移o值
输出0-6V测试结果 | 0.0058533 | -0.0067606 | 1.002 | 1.155 |  | 校准0-6V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差
0-6V(0-1023) | 0 | 0 | 0.0014 | 0.0014 | 0.02333% | 0 | 0 | 0 | 0%
| 171 | 1 | 0.9928 | -0.0072 | -0.12% | 1 | 0.9986 | -0.0014 | -0.02333% | SLOPE 函数根据提供的 y 值和 x 值返回回归线的斜率 INTERCEPT 函数用于计算直线与 y 轴的交点，基于已知的 x 值和 y 值绘制最佳拟合回归线
| 341 | 2 | 1.9923 | -0.0077 | -0.1283% | 2 | 1.9982 | -0.0018 | -0.03%
| 512 | 3 | 2.9867 | -0.0133 | -0.2217% | 3 | 2.9984 | -0.0016 | -0.02667%
| 682 | 4 | 3.9882 | -0.0118 | -0.1967% | 4 | 3.9998 | -0.0002 | -0.003333%
| 853 | 5 | 4.9848 | -0.0152 | -0.2533% | 5 | 4.9965 | -0.0035 | -0.05833%
| 1023 | 6 | 5.9942 | -0.0058 | -0.09667% | 6 | 5.9942 | -0.0058 | -0.09667%
为什么需要拟合校准？ 答：因为存在系统存在一定误差，可以通过一定的办法来减小这个误差
| 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-12V测试结果 | 0.011716 | -0.011073 | 1.0012 | 0.94515 |  | 校准0-12V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差
0-12V(0-1023) | 0 | 0 | 0.0028 | 0.0028 | - | 0 | 0.0026 | 0.0026 | 0.02167%
| 85 | 1 | 0.9936 | -0.0064 | -0.05333% | 1 | 1.005 | 0.005 | 0.04167%
| 171 | 2 | 1.9873 | -0.0127 | -0.1058% | 2 | 1.9987 | -0.0013 | -0.01083%
| 256 | 3 | 2.9813 | -0.0187 | -0.1558% | 3 | 2.9928 | -0.0072 | -0.06%
| 341 | 4 | 3.9882 | -0.0118 | -0.09833% | 4 | 3.9992 | -0.0008 | -0.006667%
| 426 | 5 | 4.9863 | -0.0137 | -0.1142% | 5 | 4.9946 | -0.0054 | -0.045%
| 512 | 6 | 5.9791 | -0.0209 | -0.1742% | 6 | 5.9905 | -0.0095 | -0.07917%
| 597 | 7 | 6.9742 | -0.0258 | -0.215% | 7 | 6.9871 | -0.0129 | -0.1075%
| 682 | 8 | 7.984 | -0.016 | -0.1333% | 8 | 7.9968 | -0.0032 | -0.02667%
| 767 | 9 | 8.9823 | -0.0177 | -0.1475% | 9 | 8.9935 | -0.0065 | -0.05417%
| 853 | 10 | 9.9803 | -0.0197 | -0.1642% | 10 | 9.9918 | -0.0082 | -0.06833%
| 938 | 11 | 10.9798 | -0.0202 | -0.1683% | 11 | 11.002 | 0.002 | 0.01667%
| 1023 | 12 | 12.006 | 0.006 | 0.05% | 12 | 12.0015 | 0.0015 | 0.0125%
|  | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
|  | 0.01955 | -0.03235 | 1 | 1.6547
输出0-20mA测试结果 | [5 empty] | 校准0-24mA | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差
0-20MA(0-1023) | 0 | 0 | 0.0012 | 0.0012 | 0.006% | 0 | 0.007 | 0.007 | 0.02917%
| 102 | 2 | 1.9826 | -0.0174 | -0.087% | 2 | 1.9804 | -0.0196 | -0.08167%
| 205 | 4 | 3.9704 | -0.0296 | -0.148% | 4 | 3.9545 | -0.0455 | -0.1896%
| 307 | 6 | 5.9584 | -0.0416 | -0.208% | 6 | 5.9539 | -0.0461 | -0.1921%
| 409 | 8 | 7.9683 | -0.0317 | -0.1585% | 8 | 7.9771 | -0.0229 | -0.09542%
| 512 | 10 | 9.9606 | -0.0394 | -0.197% | 10 | 9.971 | -0.029 | -0.1208%
| 614 | 12 | 11.9556 | -0.0444 | -0.222% | 12 | 11.9743 | -0.0257 | -0.1071%
| 716 | 14 | 13.9726 | -0.0274 | -0.137% | 14 | 13.9673 | -0.0327 | -0.1363%
| 818 | 16 | 15.9732 | -0.0268 | -0.134% | 16 | 15.9898 | -0.0102 | -0.0425%
| 921 | 18 | 17.9765 | -0.0235 | -0.1175% | 18 | 17.9776 | -0.0224 | -0.09333%
| 1023 | 20 | 20.021 | 0.021 | 0.105% | 20 | 19.9809 | -0.0191 | -0.07958%

## GP8600-PWM-EN
Accuracy calculation of DAC module
GP8600 module | [5 empty] | fitting calibration algorithm y=c*x+o" Theoretical program parameter x, transformed program parameter y Transform coefficient c, transform offset o value |  |  |  | Calculation steps 1. calculate the slope k = slope (d11: d15, b11: b15) 2. Intercept b: =INTERCEPT(D11:D15,B11:B15) (Steps 1 and 2 omit the first and last endpoint data, because the fitting effect of other data will be affected when the endpoint data error is too large.) 3. Calculate the transformation coefficient c: = (theoretical maximum voltage/maximum program parameter)/k. 4. calculate the transformation offset o value o: =-b/k. 5. Calculate the transformed program parameter y=c*x+o
Reference datesheet parameters: 0.2% full-scale FSR, current/voltage linearity error 0.05%typ
Calibration note: the calibration needs to be matched with MCU, so the calibration data and errors here are only for reference, and the DAC module needs to be calibrated by customers themselves
PWM input | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-6V test results | 0.0058533 | -0.0067606 | 1.002 | 1.155 |  | Calibration 0-6V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-6V(0-1023) | 0 | 0 | 0.0014 | 0.0014 | 0.02333% | 0 | 0 | 0 | 0%
| 171 | 1 | 0.9928 | -0.0072 | -0.12% | 1 | 0.9986 | -0.0014 | -0.02333% | The SLOPE function returns the slope of the regression line based on the supplied y and x values. INTERCEPT function is used to calculate the intersection of straight line and Y axis, and draw the best fitting regression line based on the known X value and Y value.
| 341 | 2 | 1.9923 | -0.0077 | -0.1283% | 2 | 1.9982 | -0.0018 | -0.03%
| 512 | 3 | 2.9867 | -0.0133 | -0.2217% | 3 | 2.9984 | -0.0016 | -0.02667%
| 682 | 4 | 3.9882 | -0.0118 | -0.1967% | 4 | 3.9998 | -0.0002 | -0.003333%
| 853 | 5 | 4.9848 | -0.0152 | -0.2533% | 5 | 4.9965 | -0.0035 | -0.05833%
| 1023 | 6 | 5.9942 | -0.0058 | -0.09667% | 6 | 5.9942 | -0.0058 | -0.09667%
Why do you need fitting calibration? Answer: Because there is a certain error in the system, we can reduce this error through certain methods.
| Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-12V test results | 0.011716 | -0.011073 | 1.0012 | 0.94515 |  | Calibration 0-12V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-12V(0-1023) | 0 | 0 | 0.0028 | 0.0028 | - | 0 | 0.0026 | 0.0026 | 0.02167%
| 85 | 1 | 0.9936 | -0.0064 | -0.05333% | 1 | 1.005 | 0.005 | 0.04167%
| 171 | 2 | 1.9873 | -0.0127 | -0.1058% | 2 | 1.9987 | -0.0013 | -0.01083%
| 256 | 3 | 2.9813 | -0.0187 | -0.1558% | 3 | 2.9928 | -0.0072 | -0.06%
| 341 | 4 | 3.9882 | -0.0118 | -0.09833% | 4 | 3.9992 | -0.0008 | -0.006667%
| 426 | 5 | 4.9863 | -0.0137 | -0.1142% | 5 | 4.9946 | -0.0054 | -0.045%
| 512 | 6 | 5.9791 | -0.0209 | -0.1742% | 6 | 5.9905 | -0.0095 | -0.07917%
| 597 | 7 | 6.9742 | -0.0258 | -0.215% | 7 | 6.9871 | -0.0129 | -0.1075%
| 682 | 8 | 7.984 | -0.016 | -0.1333% | 8 | 7.9968 | -0.0032 | -0.02667%
| 767 | 9 | 8.9823 | -0.0177 | -0.1475% | 9 | 8.9935 | -0.0065 | -0.05417%
| 853 | 10 | 9.9803 | -0.0197 | -0.1642% | 10 | 9.9918 | -0.0082 | -0.06833%
| 938 | 11 | 10.9798 | -0.0202 | -0.1683% | 11 | 11.002 | 0.002 | 0.01667%
| 1023 | 12 | 12.006 | 0.006 | 0.05% | 12 | 12.0015 | 0.0015 | 0.0125%
|  | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
|  | 0.01955 | -0.03235 | 1 | 1.6547
Output 0-20mA test results | [5 empty] | Calibration 0-24mA | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-20MA(0-1023) | 0 | 0 | 0.0012 | 0.0012 | 0.006% | 0 | 0.007 | 0.007 | 0.02917%
| 102 | 2 | 1.9826 | -0.0174 | -0.087% | 2 | 1.9804 | -0.0196 | -0.08167%
| 205 | 4 | 3.9704 | -0.0296 | -0.148% | 4 | 3.9545 | -0.0455 | -0.1896%
| 307 | 6 | 5.9584 | -0.0416 | -0.208% | 6 | 5.9539 | -0.0461 | -0.1921%
| 409 | 8 | 7.9683 | -0.0317 | -0.1585% | 8 | 7.9771 | -0.0229 | -0.09542%
| 512 | 10 | 9.9606 | -0.0394 | -0.197% | 10 | 9.971 | -0.029 | -0.1208%
| 614 | 12 | 11.9556 | -0.0444 | -0.222% | 12 | 11.9743 | -0.0257 | -0.1071%
| 716 | 14 | 13.9726 | -0.0274 | -0.137% | 14 | 13.9673 | -0.0327 | -0.1363%
| 818 | 16 | 15.9732 | -0.0268 | -0.134% | 16 | 15.9898 | -0.0102 | -0.0425%
| 921 | 18 | 17.9765 | -0.0235 | -0.1175% | 18 | 17.9776 | -0.0224 | -0.09333%
| 1023 | 20 | 20.021 | 0.021 | 0.105% | 20 | 19.9809 | -0.0191 | -0.07958%

## GP8600-I2C-CN
DAC模组精度计算
GP8600模块
y=x*1.001+31.1246
参考datesheet参数：±0.2%满量程FSR，电流/电压线性度误差0.05%typ
校准说明：校准需要搭配MCU，所以这里的校准数据及误差只做参考，DAC模块需要客户自行进行校准。
I2C输入 |  | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-6V测试结果 |  | 9.1458e-05 | -0.0028466 | 1.001 | 31.1246 |  | 校准0-6V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA | 满量程误差
0-6V（0-65535） | 0 | 0 | 0.0002 | 0.0002 | 0.003333% | 0 | 0 | 0.0002 | 0.0002 | 0.003333%
| 5461 | 0.5 | 0.4983 | -0.0017 | -0.02833% | 5461 | 0.5 | 0.5016 | 0.0016 | 0.02667%
| 10923 | 1 | 0.9967 | -0.0033 | -0.055% | 10923 | 1 | 1.0006 | 0.0006 | 0.01%
| 21845 | 2 | 1.9943 | -0.0057 | -0.095% | 21845 | 2 | 1.9992 | -0.0008 | -0.01333%
| 32768 | 3 | 2.9925 | -0.0075 | -0.125% | 32768 | 3 | 2.9985 | -0.0015 | -0.025%
| 36044 | 3.3 | 3.2922 | -0.0078 | -0.13% | 36044 | 3.3 | 3.2985 | -0.0015 | -0.025%
| 43690 | 4 | 3.992 | -0.008 | -0.1333% | 43690 | 4 | 3.9991 | -0.0009 | -0.015%
| 54613 | 5 | 4.9926 | -0.0074 | -0.1233% | 54613 | 5 | 5.0009 | 0.0009 | 0.015%
| 60074 | 5.5 | 5.4933 | -0.0067 | -0.1117% | 60074 | 5.5 | 5.5021 | 0.0021 | 0.035%
| 65535 | 6 | 5.9941 | -0.0059 | -0.09833% | 65535 | 6 | 5.9941 | -0.0059 | -0.09833%
|  | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-5V测试结果 |  | 7.6164e-05 | -0.0016077 | 1.0017 | 21.1089 |  | 校准0-5V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA | 满量程误差
0-5V（0-65535） | 0 | 0 | 0.0001 | 0.0001 | 0.002% | 0 | 0 | 0.0001 | 0.0001 | 0.002%
| 6553.5 | 0.5 | 0.4983 | -0.0017 | -0.034% | 6553.5 | 0.5 | 0.5006 | 0.0006 | 0.012%
| 13107 | 1 | 0.9966 | -0.0034 | -0.068% | 13107 | 1 | 0.9997 | -0.0003 | -0.006%
| 26214 | 2 | 1.9941 | -0.0059 | -0.118% | 26214 | 2 | 1.999 | -0.001 | -0.02%
| 39321 | 3 | 2.9926 | -0.0074 | -0.148% | 39321 | 3 | 2.9993 | -0.0007 | -0.014%
| 43253.1 | 3.3 | 3.2925 | -0.0075 | -0.15% | 43253.1 | 3.3 | 3.2997 | -0.0003 | -0.006%
| 52428 | 4 | 3.9925 | -0.0075 | -0.15% | 52428 | 4 | 4.0011 | 0.0011 | 0.022%
| 65535 | 5 | 4.9936 | -0.0064 | -0.128% | 65535 | 5 | 4.9936 | -0.0064 | -0.128%
|  | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-12V测试结果 |  | 0.00018314 | -0.0075327 | 0.99982 | 41.1307 |  | 校准0-12V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA | 满量程误差
0-12V（0-65535） | 0 | 0 | 0.003 | 0.003 | 0.025% | 0 | 0 | 0.003 | 0.003 | 0.025%
| 5461.25 | 1 | 0.9973 | -0.0027 | -0.0225% | 5461.25 | 1 | 1.0044 | 0.0044 | 0.03667%
| 10922.5 | 2 | 1.9952 | -0.0048 | -0.04% | 10922.5 | 2 | 2.0019 | 0.0019 | 0.01583%
| 16383.75 | 3 | 2.993 | -0.007 | -0.05833% | 16383.75 | 3 | 2.9996 | -0.0004 | -0.003333%
| 18022.125 | 3.3 | 3.2925 | -0.0075 | -0.0625% | 18022.125 | 3.3 | 3.2991 | -0.0009 | -0.0075%
| 21845 | 4 | 3.9917 | -0.0083 | -0.06917% | 21845 | 4 | 3.9982 | -0.0018 | -0.015%
| 27306.25 | 5 | 4.991 | -0.009 | -0.075% | 27306.25 | 5 | 4.9973 | -0.0027 | -0.0225%
| 30036.875 | 5.5 | 5.491 | -0.009 | -0.075% | 30036.875 | 5.5 | 5.4974 | -0.0026 | -0.02167%
| 32767.5 | 6 | 5.9911 | -0.0089 | -0.07417% | 32767.5 | 6 | 5.9973 | -0.0027 | -0.0225%
| 38228.75 | 7 | 6.9914 | -0.0086 | -0.07167% | 38228.75 | 7 | 6.9976 | -0.0024 | -0.02%
| 43690 | 8 | 7.9924 | -0.0076 | -0.06333% | 43690 | 8 | 7.9984 | -0.0016 | -0.01333%
| 49151.25 | 9 | 8.9939 | -0.0061 | -0.05083% | 49151.25 | 9 | 8.9998 | -0.0002 | -0.001667%
| 54612.5 | 10 | 9.9963 | -0.0037 | -0.03083% | 54612.5 | 10 | 10.0018 | 0.0018 | 0.015%
| 60073.75 | 11 | 10.999 | -0.001 | -0.008333% | 60073.75 | 11 | 11.0046 | 0.0046 | 0.03833%
| 65535 | 12 | 12.0023 | 0.0023 | 0.01917% | 65535 | 12 | 12.001 | 0.001 | 0.008333%
|  | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-10V测试结果 |  | 0.00015264 | -0.0062723 | 0.9997 | 41.093 |  | 校准0-10V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA | 满量程误差
0-10V（0-65535） | 0 | 0 | 0.0026 | 0.0026 | 0.02167% | 0 | 0 | 0.0026 | 0.0026 | 0.02167%
| 6553.5 | 1 | 0.9976 | -0.0024 | -0.02% | 6553.5 | 1 | 1.0035 | 0.0035 | 0.02917%
| 13107 | 2 | 1.9955 | -0.0045 | -0.0375% | 13107 | 2 | 2.0012 | 0.0012 | 0.01%
| 19660.5 | 3 | 2.9941 | -0.0059 | -0.04917% | 19660.5 | 3 | 2.9994 | -0.0006 | -0.005%
| 21626.55 | 3.3 | 3.2938 | -0.0062 | -0.05167% | 21626.55 | 3.3 | 3.299 | -0.001 | -0.008333%
| 26214 | 4 | 3.9933 | -0.0067 | -0.05583% | 26214 | 4 | 3.9983 | -0.0017 | -0.01417%
| 32767.5 | 5 | 4.9933 | -0.0067 | -0.05583% | 32767.5 | 5 | 4.9981 | -0.0019 | -0.01583%
| 36044.25 | 5.5 | 5.4934 | -0.0066 | -0.055% | 36044.25 | 5.5 | 5.4981 | -0.0019 | -0.01583%
| 39321 | 6 | 5.9939 | -0.0061 | -0.05083% | 39321 | 6 | 5.9984 | -0.0016 | -0.01333%
| 45874.5 | 7 | 6.9953 | -0.0047 | -0.03917% | 45874.5 | 7 | 6.9994 | -0.0006 | -0.005%
| 52428 | 8 | 7.9972 | -0.0028 | -0.02333% | 52428 | 8 | 8.001 | 0.001 | 0.008333%
| 58981.5 | 9 | 9 | 0 | 0% | 58981.5 | 9 | 9.0034 | 0.0034 | 0.02833%
| 65535 | 10 | 10.003 | 0.003 | 0.025% | 65535 | 10 | 10.001 | 0.001 | 0.008333%
|  | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-20mA测试结果 |  | 0.00030551 | -0.02152 | 0.99892 | 70.4414 |  | 校准0-20mA | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA | 满量程误差
0-20MA（0-65535） | 0 | 0 | 0.003 | 0.003 | 0.015% | 0 | 0 | 0.003 | 0.003 | 0.025%
| 6554 | 2 | 1.9897 | -0.0103 | -0.0515% | 6554 | 2 | 2.0094 | 0.0094 | 0.07833%
| 13107 | 4 | 3.9857 | -0.0143 | -0.0715% | 13107 | 4 | 4.0011 | 0.0011 | 0.009167%
| 26214 | 8 | 7.9799 | -0.0201 | -0.1005% | 26214 | 8 | 7.9913 | -0.0087 | -0.0725%
| 32768 | 10 | 9.9807 | -0.0193 | -0.0965% | 32768 | 10 | 9.9893 | -0.0107 | -0.08917%
| 39321 | 12 | 11.9839 | -0.0161 | -0.0805% | 39321 | 12 | 11.9874 | -0.0126 | -0.105%
| 52428 | 16 | 15.9973 | -0.0027 | -0.0135% | 52428 | 16 | 16.003 | 0.003 | 0.025%
| 58982 | 18 | 18.008 | 0.008 | 0.04% | 58982 | 18 | 18.0103 | 0.0103 | 0.08583%
| 65535 | 20 | 20.0212 | 0.0212 | 0.106% | 65535 | 20 | 20.0212 | 0.0212 | 0.1767%
tips | 电流测量时输出时会向上飘动一会儿才稳定下来
(如开始输出18.006mA 会缓慢上升到18.014mA， 此时切换输出会从2.015mA缓慢跌落至2.0094mA)

## GP8600-I2C-EN
Accuracy calculation of DAC module
GP8600 module
y=x*1.001+31.1246
Reference datesheet parameters: 0.2% full-scale FSR, current/voltage linearity error 0.05%typ
Calibration note: the calibration needs to be matched with MCU, so the calibration data and errors here are only for reference, and the DAC module needs to be calibrated by customers themselves
I2C input |  | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-6V test results |  | 9.1458e-05 | -0.0028466 | 1.001 | 31.1246 |  | Calibration 0-6V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-6V（0-65535） | 0 | 0 | 0.0002 | 0.0002 | 0.003333% | 0 | 0 | 0.0002 | 0.0002 | 0.003333%
| 5461 | 0.5 | 0.4983 | -0.0017 | -0.02833% | 5461 | 0.5 | 0.5016 | 0.0016 | 0.02667%
| 10923 | 1 | 0.9967 | -0.0033 | -0.055% | 10923 | 1 | 1.0006 | 0.0006 | 0.01%
| 21845 | 2 | 1.9943 | -0.0057 | -0.095% | 21845 | 2 | 1.9992 | -0.0008 | -0.01333%
| 32768 | 3 | 2.9925 | -0.0075 | -0.125% | 32768 | 3 | 2.9985 | -0.0015 | -0.025%
| 36044 | 3.3 | 3.2922 | -0.0078 | -0.13% | 36044 | 3.3 | 3.2985 | -0.0015 | -0.025%
| 43690 | 4 | 3.992 | -0.008 | -0.1333% | 43690 | 4 | 3.9991 | -0.0009 | -0.015%
| 54613 | 5 | 4.9926 | -0.0074 | -0.1233% | 54613 | 5 | 5.0009 | 0.0009 | 0.015%
| 60074 | 5.5 | 5.4933 | -0.0067 | -0.1117% | 60074 | 5.5 | 5.5021 | 0.0021 | 0.035%
| 65535 | 6 | 5.9941 | -0.0059 | -0.09833% | 65535 | 6 | 5.9941 | -0.0059 | -0.09833%
|  | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-5V test results |  | 7.6164e-05 | -0.0016077 | 1.0017 | 21.1089 |  | Calibration 0-5V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-5V（0-65535） | 0 | 0 | 0.0001 | 0.0001 | 0.002% | 0 | 0 | 0.0001 | 0.0001 | 0.002%
| 6553.5 | 0.5 | 0.4983 | -0.0017 | -0.034% | 6553.5 | 0.5 | 0.5006 | 0.0006 | 0.012%
| 13107 | 1 | 0.9966 | -0.0034 | -0.068% | 13107 | 1 | 0.9997 | -0.0003 | -0.006%
| 26214 | 2 | 1.9941 | -0.0059 | -0.118% | 26214 | 2 | 1.999 | -0.001 | -0.02%
| 39321 | 3 | 2.9926 | -0.0074 | -0.148% | 39321 | 3 | 2.9993 | -0.0007 | -0.014%
| 43253.1 | 3.3 | 3.2925 | -0.0075 | -0.15% | 43253.1 | 3.3 | 3.2997 | -0.0003 | -0.006%
| 52428 | 4 | 3.9925 | -0.0075 | -0.15% | 52428 | 4 | 4.0011 | 0.0011 | 0.022%
| 65535 | 5 | 4.9936 | -0.0064 | -0.128% | 65535 | 5 | 4.9936 | -0.0064 | -0.128%
|  | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-12V test results |  | 0.00018314 | -0.0075327 | 0.99982 | 41.1307 |  | Calibration 0-12V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-12V（0-65535） | 0 | 0 | 0.003 | 0.003 | 0.025% | 0 | 0 | 0.003 | 0.003 | 0.025%
| 5461.25 | 1 | 0.9973 | -0.0027 | -0.0225% | 5461.25 | 1 | 1.0044 | 0.0044 | 0.03667%
| 10922.5 | 2 | 1.9952 | -0.0048 | -0.04% | 10922.5 | 2 | 2.0019 | 0.0019 | 0.01583%
| 16383.75 | 3 | 2.993 | -0.007 | -0.05833% | 16383.75 | 3 | 2.9996 | -0.0004 | -0.003333%
| 18022.125 | 3.3 | 3.2925 | -0.0075 | -0.0625% | 18022.125 | 3.3 | 3.2991 | -0.0009 | -0.0075%
| 21845 | 4 | 3.9917 | -0.0083 | -0.06917% | 21845 | 4 | 3.9982 | -0.0018 | -0.015%
| 27306.25 | 5 | 4.991 | -0.009 | -0.075% | 27306.25 | 5 | 4.9973 | -0.0027 | -0.0225%
| 30036.875 | 5.5 | 5.491 | -0.009 | -0.075% | 30036.875 | 5.5 | 5.4974 | -0.0026 | -0.02167%
| 32767.5 | 6 | 5.9911 | -0.0089 | -0.07417% | 32767.5 | 6 | 5.9973 | -0.0027 | -0.0225%
| 38228.75 | 7 | 6.9914 | -0.0086 | -0.07167% | 38228.75 | 7 | 6.9976 | -0.0024 | -0.02%
| 43690 | 8 | 7.9924 | -0.0076 | -0.06333% | 43690 | 8 | 7.9984 | -0.0016 | -0.01333%
| 49151.25 | 9 | 8.9939 | -0.0061 | -0.05083% | 49151.25 | 9 | 8.9998 | -0.0002 | -0.001667%
| 54612.5 | 10 | 9.9963 | -0.0037 | -0.03083% | 54612.5 | 10 | 10.0018 | 0.0018 | 0.015%
| 60073.75 | 11 | 10.999 | -0.001 | -0.008333% | 60073.75 | 11 | 11.0046 | 0.0046 | 0.03833%
| 65535 | 12 | 12.0023 | 0.0023 | 0.01917% | 65535 | 12 | 12.001 | 0.001 | 0.008333%
|  | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-10V test results |  | 0.00015264 | -0.0062723 | 0.9997 | 41.093 |  | Calibration 0-10V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-10V（0-65535） | 0 | 0 | 0.0026 | 0.0026 | 0.02167% | 0 | 0 | 0.0026 | 0.0026 | 0.02167%
| 6553.5 | 1 | 0.9976 | -0.0024 | -0.02% | 6553.5 | 1 | 1.0035 | 0.0035 | 0.02917%
| 13107 | 2 | 1.9955 | -0.0045 | -0.0375% | 13107 | 2 | 2.0012 | 0.0012 | 0.01%
| 19660.5 | 3 | 2.9941 | -0.0059 | -0.04917% | 19660.5 | 3 | 2.9994 | -0.0006 | -0.005%
| 21626.55 | 3.3 | 3.2938 | -0.0062 | -0.05167% | 21626.55 | 3.3 | 3.299 | -0.001 | -0.008333%
| 26214 | 4 | 3.9933 | -0.0067 | -0.05583% | 26214 | 4 | 3.9983 | -0.0017 | -0.01417%
| 32767.5 | 5 | 4.9933 | -0.0067 | -0.05583% | 32767.5 | 5 | 4.9981 | -0.0019 | -0.01583%
| 36044.25 | 5.5 | 5.4934 | -0.0066 | -0.055% | 36044.25 | 5.5 | 5.4981 | -0.0019 | -0.01583%
| 39321 | 6 | 5.9939 | -0.0061 | -0.05083% | 39321 | 6 | 5.9984 | -0.0016 | -0.01333%
| 45874.5 | 7 | 6.9953 | -0.0047 | -0.03917% | 45874.5 | 7 | 6.9994 | -0.0006 | -0.005%
| 52428 | 8 | 7.9972 | -0.0028 | -0.02333% | 52428 | 8 | 8.001 | 0.001 | 0.008333%
| 58981.5 | 9 | 9 | 0 | 0% | 58981.5 | 9 | 9.0034 | 0.0034 | 0.02833%
| 65535 | 10 | 10.003 | 0.003 | 0.025% | 65535 | 10 | 10.001 | 0.001 | 0.008333%
|  | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-20mA test results |  | 0.00030551 | -0.02152 | 0.99892 | 70.4414 |  | Calibration 0-20mA | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-20MA（0-65535） | 0 | 0 | 0.003 | 0.003 | 0.015% | 0 | 0 | 0.003 | 0.003 | 0.025%
| 6554 | 2 | 1.9897 | -0.0103 | -0.0515% | 6554 | 2 | 2.0094 | 0.0094 | 0.07833%
| 13107 | 4 | 3.9857 | -0.0143 | -0.0715% | 13107 | 4 | 4.0011 | 0.0011 | 0.009167%
| 26214 | 8 | 7.9799 | -0.0201 | -0.1005% | 26214 | 8 | 7.9913 | -0.0087 | -0.0725%
| 32768 | 10 | 9.9807 | -0.0193 | -0.0965% | 32768 | 10 | 9.9893 | -0.0107 | -0.08917%
| 39321 | 12 | 11.9839 | -0.0161 | -0.0805% | 39321 | 12 | 11.9874 | -0.0126 | -0.105%
| 52428 | 16 | 15.9973 | -0.0027 | -0.0135% | 52428 | 16 | 16.003 | 0.003 | 0.025%
| 58982 | 18 | 18.008 | 0.008 | 0.04% | 58982 | 18 | 18.0103 | 0.0103 | 0.08583%
| 65535 | 20 | 20.0212 | 0.0212 | 0.106% | 65535 | 20 | 20.0212 | 0.0212 | 0.1767%
tips | When the current is measured, the output will float upward for a while before it stabilizes.
(If the output starts at 18.006mA, it will slowly rise to 18.014mA, At this time, the switching output will slowly drop from 2.015mA to 2.0094mA)

## GP8630N-PWM-CN
DAC模组精度计算
GP8630N模块
参考datesheet参数：输出误差（满量程偏差）±0.1%，电压线性度0.01%typ，电流线性度0.02%typ
校准说明：校准需要搭配MCU，所以这里的校准数据及误差只做参考，DAC模块需要客户自行进行校准。
PWM输入 | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-12V测试结果 | 0.01171 | 0.001082 | 1.0018 | -0.092403 |  | 校准0-12V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 绝对误差 | 满量程误差
0-12V（0-1023） | 0 | 0 | 0.0052 | 0.0052 | 0.04333% | 0 | 0.0052 | 0.0052 | 0.04333%
| 85 | 1 | 1.002 | 0.002 | 0.01667% | 1 | 1.0025 | 0.0025 | 0.02083%
| 171 | 2 | 1.9961 | -0.0039 | -0.0325% | 2 | 1.996 | -0.004 | -0.03333%
| 256 | 3 | 2.9926 | -0.0074 | -0.06167% | 3 | 2.9926 | -0.0074 | -0.06167%
| 341 | 4 | 4.0012 | 0.0012 | 0.01% | 4 | 4.0011 | 0.0011 | 0.009167%
| 426 | 5 | 4.9951 | -0.0049 | -0.04083% | 5 | 4.995 | -0.005 | -0.04167%
| 512 | 6 | 5.9924 | -0.0076 | -0.06333% | 6 | 5.9923 | -0.0077 | -0.06417%
| 597 | 7 | 6.9846 | -0.0154 | -0.1283% | 7 | 6.9845 | -0.0155 | -0.1292%
| 682 | 8 | 7.9912 | -0.0088 | -0.07333% | 8 | 7.9911 | -0.0089 | -0.07417%
| 767 | 9 | 8.9891 | -0.0109 | -0.09083% | 9 | 9.0006 | 0.0006 | 0.005%
| 853 | 10 | 9.9868 | -0.0132 | -0.11% | 10 | 9.9986 | -0.0014 | -0.01167%
| 938 | 11 | 10.9826 | -0.0174 | -0.145% | 11 | 10.9943 | -0.0057 | -0.0475%
| 1023 | 12 | 12.006 | 0.006 | 0.05% | 12 | 12.0059 | 0.0059 | 0.04917%
| 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出-12-0V测试结果 | -0.011714 | -0.0018087 | 1.0013 | -0.1544 |  | 校准-12-0V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 绝对误差 | 满量程误差
-12-0V（1023-0） | 0 | 0 | -0.0007 | -0.0007 | 0.005833% | 0 | -0.0007 | -0.0007 | -0.005833%
| 85 | -1 | -1.0005 | -0.0005 | 0.004167% | -1 | -0.9885 | 0.0115 | 0.09583%
| 171 | -2 | -1.9959 | 0.0041 | -0.03417% | -2 | -1.9959 | 0.0041 | 0.03417%
| 256 | -3 | -2.994 | 0.006 | -0.05% | -3 | -2.994 | 0.006 | 0.05%
| 341 | -4 | -4.0037 | -0.0037 | 0.03083% | -4 | -4.0037 | -0.0037 | -0.03083%
| 426 | -5 | -4.9986 | 0.0014 | -0.01167% | -5 | -4.9986 | 0.0014 | 0.01167%
| 512 | -6 | -5.9963 | 0.0037 | -0.03083% | -6 | -5.9963 | 0.0037 | 0.03083%
| 597 | -7 | -6.9988 | 0.0012 | -0.01% | -7 | -6.9885 | 0.0115 | 0.09583%
| 682 | -8 | -7.9953 | 0.0047 | -0.03917% | -8 | -7.9953 | 0.0047 | 0.03917%
| 767 | -9 | -8.9926 | 0.0074 | -0.06167% | -9 | -8.9921 | 0.0079 | 0.06583%
| 853 | -10 | -9.9893 | 0.0107 | -0.08917% | -10 | -9.9893 | 0.0107 | 0.08917%
| 938 | -11 | -10.984 | 0.016 | -0.1333% | -11 | -10.995 | 0.005 | 0.04167%
| 1023 | -12 | -12.0056 | -0.0056 | 0.04667% | -12 | -12.0056 | -0.0056 | -0.04667%
| 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
| 0.023452 | 0.039402 | 1.0004 | -1.6801
输出0-24mA测试结果 | [5 empty] | 校准0-24mA | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 绝对误差 | 满量程误差
0-24MA（0-1023） | 0 | 0 | 0.007 | 0.007 | 0.02917% | 0 | 0.007 | 0.007 | 0.02917%
| 85.25 | 2 | 2.1372 | 0.1372 | 0.5717% | 2 | 1.9804 | -0.0196 | -0.08167%
| 170.5 | 4 | 4.0013 | 0.0013 | 0.005417% | 4 | 3.9545 | -0.0455 | -0.1896%
| 255.75 | 6 | 6.0002 | 0.0002 | 0.0008333% | 6 | 5.9539 | -0.0461 | -0.1921%
| 341 | 8 | 8.0242 | 0.0242 | 0.1008% | 8 | 7.9771 | -0.0229 | -0.09542%
| 426.25 | 10 | 10.018 | 0.018 | 0.075% | 10 | 9.971 | -0.029 | -0.1208%
| 511.5 | 12 | 12.0207 | 0.0207 | 0.08625% | 12 | 11.9743 | -0.0257 | -0.1071%
| 596.75 | 14 | 14.0134 | 0.0134 | 0.05583% | 14 | 13.9673 | -0.0327 | -0.1363%
| 682 | 16 | 16.0138 | 0.0138 | 0.0575% | 16 | 15.9898 | -0.0102 | -0.0425%
| 767.25 | 18 | 18.0438 | 0.0438 | 0.1825% | 18 | 17.9776 | -0.0224 | -0.09333%
| 852.5 | 20 | 20.0529 | 0.0529 | 0.2204% | 20 | 19.9809 | -0.0191 | -0.07958%
| 937.75 | 22 | 22.0583 | 0.0583 | 0.2429% | 22 | 21.9612 | -0.0388 | -0.1617%
| 1023 | 24 | 24.12 | 0.12 | 0.5% | 24 | 24.05 | 0.05 | 0.2083%

## GP8630N-PWM-EN
Accuracy calculation of DAC module
GP8630N module
Refer to datesheet parameters: output error (full-scale deviation) 0.1%, voltage linearity 0.01%typ, current linearity 0.02%typ
Calibration note: the calibration needs to be matched with MCU, so the calibration data and errors here are only for reference, and the DAC module needs to be calibrated by customers themselves
PWM input | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-12V test results | 0.01171 | 0.001082 | 1.0018 | -0.092403 |  | Calibration 0-12V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-12V（0-1023） | 0 | 0 | 0.0052 | 0.0052 | 0.04333% | 0 | 0.0052 | 0.0052 | 0.04333%
| 85 | 1 | 1.002 | 0.002 | 0.01667% | 1 | 1.0025 | 0.0025 | 0.02083%
| 171 | 2 | 1.9961 | -0.0039 | -0.0325% | 2 | 1.996 | -0.004 | -0.03333%
| 256 | 3 | 2.9926 | -0.0074 | -0.06167% | 3 | 2.9926 | -0.0074 | -0.06167%
| 341 | 4 | 4.0012 | 0.0012 | 0.01% | 4 | 4.0011 | 0.0011 | 0.009167%
| 426 | 5 | 4.9951 | -0.0049 | -0.04083% | 5 | 4.995 | -0.005 | -0.04167%
| 512 | 6 | 5.9924 | -0.0076 | -0.06333% | 6 | 5.9923 | -0.0077 | -0.06417%
| 597 | 7 | 6.9846 | -0.0154 | -0.1283% | 7 | 6.9845 | -0.0155 | -0.1292%
| 682 | 8 | 7.9912 | -0.0088 | -0.07333% | 8 | 7.9911 | -0.0089 | -0.07417%
| 767 | 9 | 8.9891 | -0.0109 | -0.09083% | 9 | 9.0006 | 0.0006 | 0.005%
| 853 | 10 | 9.9868 | -0.0132 | -0.11% | 10 | 9.9986 | -0.0014 | -0.01167%
| 938 | 11 | 10.9826 | -0.0174 | -0.145% | 11 | 10.9943 | -0.0057 | -0.0475%
| 1023 | 12 | 12.006 | 0.006 | 0.05% | 12 | 12.0059 | 0.0059 | 0.04917%
| Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output -12-0V test results | -0.011714 | -0.0018087 | 1.0013 | -0.1544 |  | Calibration -12-0V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
-12-0V（1023-0） | 0 | 0 | -0.0007 | -0.0007 | 0.005833% | 0 | -0.0007 | -0.0007 | -0.005833%
| 85 | -1 | -1.0005 | -0.0005 | 0.004167% | -1 | -0.9885 | 0.0115 | 0.09583%
| 171 | -2 | -1.9959 | 0.0041 | -0.03417% | -2 | -1.9959 | 0.0041 | 0.03417%
| 256 | -3 | -2.994 | 0.006 | -0.05% | -3 | -2.994 | 0.006 | 0.05%
| 341 | -4 | -4.0037 | -0.0037 | 0.03083% | -4 | -4.0037 | -0.0037 | -0.03083%
| 426 | -5 | -4.9986 | 0.0014 | -0.01167% | -5 | -4.9986 | 0.0014 | 0.01167%
| 512 | -6 | -5.9963 | 0.0037 | -0.03083% | -6 | -5.9963 | 0.0037 | 0.03083%
| 597 | -7 | -6.9988 | 0.0012 | -0.01% | -7 | -6.9885 | 0.0115 | 0.09583%
| 682 | -8 | -7.9953 | 0.0047 | -0.03917% | -8 | -7.9953 | 0.0047 | 0.03917%
| 767 | -9 | -8.9926 | 0.0074 | -0.06167% | -9 | -8.9921 | 0.0079 | 0.06583%
| 853 | -10 | -9.9893 | 0.0107 | -0.08917% | -10 | -9.9893 | 0.0107 | 0.08917%
| 938 | -11 | -10.984 | 0.016 | -0.1333% | -11 | -10.995 | 0.005 | 0.04167%
| 1023 | -12 | -12.0056 | -0.0056 | 0.04667% | -12 | -12.0056 | -0.0056 | -0.04667%
| Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
| 0.023452 | 0.039402 | 1.0004 | -1.6801
Output 0-24mA test results | [5 empty] | Calibration 0-24mA | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-24MA（0-1023） | 0 | 0 | 0.007 | 0.007 | 0.02917% | 0 | 0.007 | 0.007 | 0.02917%
| 85.25 | 2 | 2.1372 | 0.1372 | 0.5717% | 2 | 1.9804 | -0.0196 | -0.08167%
| 170.5 | 4 | 4.0013 | 0.0013 | 0.005417% | 4 | 3.9545 | -0.0455 | -0.1896%
| 255.75 | 6 | 6.0002 | 0.0002 | 0.0008333% | 6 | 5.9539 | -0.0461 | -0.1921%
| 341 | 8 | 8.0242 | 0.0242 | 0.1008% | 8 | 7.9771 | -0.0229 | -0.09542%
| 426.25 | 10 | 10.018 | 0.018 | 0.075% | 10 | 9.971 | -0.029 | -0.1208%
| 511.5 | 12 | 12.0207 | 0.0207 | 0.08625% | 12 | 11.9743 | -0.0257 | -0.1071%
| 596.75 | 14 | 14.0134 | 0.0134 | 0.05583% | 14 | 13.9673 | -0.0327 | -0.1363%
| 682 | 16 | 16.0138 | 0.0138 | 0.0575% | 16 | 15.9898 | -0.0102 | -0.0425%
| 767.25 | 18 | 18.0438 | 0.0438 | 0.1825% | 18 | 17.9776 | -0.0224 | -0.09333%
| 852.5 | 20 | 20.0529 | 0.0529 | 0.2204% | 20 | 19.9809 | -0.0191 | -0.07958%
| 937.75 | 22 | 22.0583 | 0.0583 | 0.2429% | 22 | 21.9612 | -0.0388 | -0.1617%
| 1023 | 24 | 24.12 | 0.12 | 0.5% | 24 | 24.05 | 0.05 | 0.2083%

## GP8630N-I2C-CN
DAC模组精度计算
GP8630N模块
参考datesheet参数：输出误差（满量程偏差）±0.1%，电压线性度0.01%typ，电流线性度0.02%typ
校准说明：校准需要搭配MCU，所以这里的校准数据及误差只做参考，DAC模块需要客户自行进行校准。 | [5 empty] | 此时变换系数为1，偏移为-5.208(相对于65535来说太小)， 且未校准前已经比较精确，继续校准意义不大
I2C输入 | 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-12V测试结果 | 0.00018311 | 0.00095362 | 1 | -5.208
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差
0-12V（0-65535） | 0 | 0 | 0.0049 | 0.0049 | 0.04083%
| 5461 | 1 | 0.9941 | -0.0059 | -0.04917%
| 10923 | 2 | 2.0041 | 0.0041 | 0.03417%
| 16384 | 3 | 3.0035 | 0.0035 | 0.02917%
| 21845 | 4 | 4.0052 | 0.0052 | 0.04333%
| 27306 | 5 | 5.003 | 0.003 | 0.025%
| 32768 | 6 | 5.9982 | -0.0018 | -0.015%
| 38229 | 7 | 7.0005 | 0.0005 | 0.004167%
| 43690 | 8 | 7.9988 | -0.0012 | -0.01%
| 49151 | 9 | 8.9997 | -0.0003 | -0.0025%
| 54613 | 10 | 10.0019 | 0.0019 | 0.01583%
| 60074 | 11 | 11.001 | 0.001 | 0.008333%
| 65535 | 12 | 12.0043 | 0.0043 | 0.03583%
输出0-10V测试结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差
0-10V（0-65535） | 0 | 0 | 0.005 | 0.005 | 0.04167%
| 6553.5 | 1 | 1.0046 | 0.0046 | 0.03833%
| 13107 | 2 | 2.0055 | 0.0055 | 0.04583%
| 19660.5 | 3 | 3.0004 | 0.0004 | 0.003333%
| 26214 | 4 | 4.0032 | 0.0032 | 0.02667%
| 32767.5 | 5 | 4.9993 | -0.0007 | -0.005833%
| 39321 | 6 | 6.0001 | 0.0001 | 0.0008333%
| 45874.5 | 7 | 7.0007 | 0.0007 | 0.005833%
| 52428 | 8 | 8 | 0 | 0%
| 58981.5 | 9 | 9.0003 | 0.0003 | 0.0025%
| 65535 | 10 | 10.0019 | 0.0019 | 0.01583%
| 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出-12-0V测试结果 | -0.00018306 | -0.0058009 | 1.0003 | -31.6891 |  |  | 校准-12-0V | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 |  | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差 | 满量程误差
-12-0V（65535-0） | 0 | 0 | 0.000781 | 0.000781 | -0.006508% |  | 0 | 0 | 0 | 0%
| 5461 | -1 | -1.0044 | -0.0044 | 0.03667% |  | -1 | -0.9987 | 0.0013 | -0.01083%
| 10923 | -2 | -2.0038 | -0.0038 | 0.03167% |  | -2 | -1.9986 | 0.0014 | -0.01167%
| 16384 | -3 | -3.005 | -0.005 | 0.04167% |  | -3 | -3.0009 | -0.0009 | 0.0075%
| 21845 | -4 | -4.0079 | -0.0079 | 0.06583% |  | -4 | -4.0031 | -0.0031 | 0.02583%
| 27306 | -5 | -5.0066 | -0.0066 | 0.055% |  | -5 | -5.002 | -0.002 | 0.01667%
| 32768 | -6 | -6.0025 | -0.0025 | 0.02083% |  | -6 | -6.0042 | -0.0042 | 0.035%
| 38229 | -7 | -7.0048 | -0.0048 | 0.04% |  | -7 | -7.0005 | -0.0005 | 0.004167%
| 43690 | -8 | -8.003 | -0.003 | 0.025% |  | -8 | -7.9993 | 0.0007 | -0.005833%
| 49151 | -9 | -9.0003 | -0.0003 | 0.0025% |  | -9 | -9.0003 | -0.0003 | 0.0025%
| 54613 | -10 | -10.0048 | -0.0048 | 0.04% |  | -10 | -10.0016 | -0.0016 | 0.01333%
| 60074 | -11 | -11.0027 | -0.0027 | 0.0225% |  | -11 | -10.9996 | 0.0004 | -0.003333%
| 65535 | -12 | -12.0046 | -0.0046 | 0.03833% |  | -12 | -12.0002 | -0.00018 | 0.0015%
输出-10-0V测试结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差
-10-0V（0-65535） | 0 | 0 | 0.0006 | 0.0006 | 0.005%
| 6553.5 | -1 | -1.0041 | 0.00406 | 0.03383%
| 13107 | -2 | -2.0051 | 0.0051 | 0.0425%
| 19660.5 | -3 | -3.0051 | 0.0051 | 0.0425%
| 26214 | -4 | -4.0058 | 0.0058 | 0.04833%
| 32767.5 | -5 | -5.0028 | 0.0028 | 0.02333%
| 39321 | -6 | -6.0047 | 0.0047 | 0.03917%
| 45874.5 | -7 | -7.0051 | 0.0051 | 0.0425%
| 52428 | -8 | -8.0041 | 0.0041 | 0.03417%
| 58981.5 | -9 | -9.0003 | 0.0003 | 0.0025%
| 65535 | -10 | -10.005 | 0.005 | 0.04167%
| 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-24mA测试结果 | 0.00036743 | 0.00028884 | 0.9967 | -0.78609 |  |  | 校准0-24mA | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 |  | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差 | 满量程误差
0-24mA（0-65535） | 0 | 0 | 0.0042 | 0.0042 | 0.0175% |  | 0 | 0 | 0 | 0%
| 5461 | 2 | 2.0127 | 0.0127 | 0.05292% |  | 2 | 2.0098 | 0.0098 | 0.04083%
| 10923 | 4 | 4.0147 | 0.0147 | 0.06125% |  | 4 | 4.0032 | 0.0032 | 0.01333%
| 21845 | 8 | 8.0287 | 0.0287 | 0.1196% |  | 8 | 8.0042 | 0.0042 | 0.0175%
| 27306 | 10 | 10.0311 | 0.0311 | 0.1296% |  | 10 | 10 | 0 | 0%
| 32768 | 12 | 12.0279 | 0.0279 | 0.1163% |  | 12 | 12.0032 | 0.0032 | 0.01333%
| 43690 | 16 | 16.0476 | 0.0476 | 0.1983% |  | 16 | 15.9957 | -0.0043 | -0.01792%
| 49151 | 18 | 18.0602 | 0.0602 | 0.2508% |  | 18 | 18.0024 | 0.0024 | 0.01%
| 54613 | 20 | 20.0781 | 0.0781 | 0.3254% |  | 20 | 20.0115 | 0.0115 | 0.04792%
| 65535 | 24 | 24.1111 | 0.1111 | 0.4629% |  | 24 | 24.033 | 0.033 | 0.1375%
| 实际斜率k | 实际截距b | 变换系数 | 变换偏移值
输出0-20mA测试结果 | 0.0003061 | 0.0044508 | 0.99699 | -14.5404 |  |  | 校准0-20mA | 拟合校准之后的数据结果
量程 | 程序参数 | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差（V/mA） | 满量程误差 |  | 理论输出电压/电流（V/mA） | 实际输出电压/电流（V/mA） | 误差 | 满量程误差
0-24mA（0-65535） | 0 | 0 | 0.0067 | 0.0067 | 0.02792% |  | 0 | 0 | 0 | 0%
| 6553.5 | 2 | 2.014 | 0.014 | 0.05833% |  | 2 | 2.0054 | 0.0054 | 0.0225%
| 13107 | 4 | 4.0197 | 0.0197 | 0.08208% |  | 4 | 4.0035 | 0.0035 | 0.01458%
| 26214 | 8 | 8.0276 | 0.0276 | 0.115% |  | 8 | 8 | 0 | 0%
| 32767.5 | 10 | 10.0259 | 0.0259 | 0.1079% |  | 10 | 10.0035 | 0.0035 | 0.01458%
| 39321 | 12 | 12.037 | 0.037 | 0.1542% |  | 12 | 11.997 | -0.003 | -0.0125%
| 52428 | 16 | 16.053 | 0.053 | 0.2208% |  | 16 | 16.0015 | 0.0015 | 0.00625%
| 58981.5 | 18 | 18.0651 | 0.0651 | 0.2713% |  | 18 | 18.0075 | 0.0075 | 0.03125%
| 65535 | 20 | 20.0813 | 0.0813 | 0.3387% |  | 20 | 20.0168 | 0.0168 | 0.07%

## GP8630N-I2C-EN
Accuracy calculation of DAC module
GP8630N module
Refer to datesheet parameters: output error (full-scale deviation) 0.1%, voltage linearity 0.01%typ, current linearity 0.02%typ
Calibration note: the calibration needs to be matched with MCU, so the calibration data and errors here are only for reference, and the DAC module needs to be calibrated by customers themselves | [5 empty] | At this time, the transform coefficient is 1 and the offset is -5.208 (too small for 65535). And it has been more accurate before calibration, so it is of little significance to continue calibration.
I2C input | Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-12V test results | 0.00018311 | 0.00095362 | 1 | -5.208
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-12V（0-65535） | 0 | 0 | 0.0049 | 0.0049 | 0.04083%
| 5461 | 1 | 0.9941 | -0.0059 | -0.04917%
| 10923 | 2 | 2.0041 | 0.0041 | 0.03417%
| 16384 | 3 | 3.0035 | 0.0035 | 0.02917%
| 21845 | 4 | 4.0052 | 0.0052 | 0.04333%
| 27306 | 5 | 5.003 | 0.003 | 0.025%
| 32768 | 6 | 5.9982 | -0.0018 | -0.015%
| 38229 | 7 | 7.0005 | 0.0005 | 0.004167%
| 43690 | 8 | 7.9988 | -0.0012 | -0.01%
| 49151 | 9 | 8.9997 | -0.0003 | -0.0025%
| 54613 | 10 | 10.0019 | 0.0019 | 0.01583%
| 60074 | 11 | 11.001 | 0.001 | 0.008333%
| 65535 | 12 | 12.0043 | 0.0043 | 0.03583%
Output 0-10V test results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-10V（0-65535） | 0 | 0 | 0.005 | 0.005 | 0.04167%
| 6553.5 | 1 | 1.0046 | 0.0046 | 0.03833%
| 13107 | 2 | 2.0055 | 0.0055 | 0.04583%
| 19660.5 | 3 | 3.0004 | 0.0004 | 0.003333%
| 26214 | 4 | 4.0032 | 0.0032 | 0.02667%
| 32767.5 | 5 | 4.9993 | -0.0007 | -0.005833%
| 39321 | 6 | 6.0001 | 0.0001 | 0.0008333%
| 45874.5 | 7 | 7.0007 | 0.0007 | 0.005833%
| 52428 | 8 | 8 | 0 | 0%
| 58981.5 | 9 | 9.0003 | 0.0003 | 0.0025%
| 65535 | 10 | 10.0019 | 0.0019 | 0.01583%
| Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output -12-0V test results | -0.00018306 | -0.0058009 | 1.0003 | -31.6891 |  |  | Calibration -12-0V | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error |  | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
-12-0V（65535-0） | 0 | 0 | 0.000781 | 0.000781 | -0.006508% |  | 0 | 0 | 0 | 0%
| 5461 | -1 | -1.0044 | -0.0044 | 0.03667% |  | -1 | -0.9987 | 0.0013 | -0.01083%
| 10923 | -2 | -2.0038 | -0.0038 | 0.03167% |  | -2 | -1.9986 | 0.0014 | -0.01167%
| 16384 | -3 | -3.005 | -0.005 | 0.04167% |  | -3 | -3.0009 | -0.0009 | 0.0075%
| 21845 | -4 | -4.0079 | -0.0079 | 0.06583% |  | -4 | -4.0031 | -0.0031 | 0.02583%
| 27306 | -5 | -5.0066 | -0.0066 | 0.055% |  | -5 | -5.002 | -0.002 | 0.01667%
| 32768 | -6 | -6.0025 | -0.0025 | 0.02083% |  | -6 | -6.0042 | -0.0042 | 0.035%
| 38229 | -7 | -7.0048 | -0.0048 | 0.04% |  | -7 | -7.0005 | -0.0005 | 0.004167%
| 43690 | -8 | -8.003 | -0.003 | 0.025% |  | -8 | -7.9993 | 0.0007 | -0.005833%
| 49151 | -9 | -9.0003 | -0.0003 | 0.0025% |  | -9 | -9.0003 | -0.0003 | 0.0025%
| 54613 | -10 | -10.0048 | -0.0048 | 0.04% |  | -10 | -10.0016 | -0.0016 | 0.01333%
| 60074 | -11 | -11.0027 | -0.0027 | 0.0225% |  | -11 | -10.9996 | 0.0004 | -0.003333%
| 65535 | -12 | -12.0046 | -0.0046 | 0.03833% |  | -12 | -12.0002 | -0.00018 | 0.0015%
Output -10-0V test results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
-10-0V（0-65535） | 0 | 0 | 0.0006 | 0.0006 | 0.005%
| 6553.5 | -1 | -1.0041 | 0.00406 | 0.03383%
| 13107 | -2 | -2.0051 | 0.0051 | 0.0425%
| 19660.5 | -3 | -3.0051 | 0.0051 | 0.0425%
| 26214 | -4 | -4.0058 | 0.0058 | 0.04833%
| 32767.5 | -5 | -5.0028 | 0.0028 | 0.02333%
| 39321 | -6 | -6.0047 | 0.0047 | 0.03917%
| 45874.5 | -7 | -7.0051 | 0.0051 | 0.0425%
| 52428 | -8 | -8.0041 | 0.0041 | 0.03417%
| 58981.5 | -9 | -9.0003 | 0.0003 | 0.0025%
| 65535 | -10 | -10.005 | 0.005 | 0.04167%
| Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-24mA test results | 0.00036743 | 0.00028884 | 0.9967 | -0.78609 |  |  | Calibration 0-24mA | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error |  | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-24mA（0-65535） | 0 | 0 | 0.0042 | 0.0042 | 0.0175% |  | 0 | 0 | 0 | 0%
| 5461 | 2 | 2.0127 | 0.0127 | 0.05292% |  | 2 | 2.0098 | 0.0098 | 0.04083%
| 10923 | 4 | 4.0147 | 0.0147 | 0.06125% |  | 4 | 4.0032 | 0.0032 | 0.01333%
| 21845 | 8 | 8.0287 | 0.0287 | 0.1196% |  | 8 | 8.0042 | 0.0042 | 0.0175%
| 27306 | 10 | 10.0311 | 0.0311 | 0.1296% |  | 10 | 10 | 0 | 0%
| 32768 | 12 | 12.0279 | 0.0279 | 0.1163% |  | 12 | 12.0032 | 0.0032 | 0.01333%
| 43690 | 16 | 16.0476 | 0.0476 | 0.1983% |  | 16 | 15.9957 | -0.0043 | -0.01792%
| 49151 | 18 | 18.0602 | 0.0602 | 0.2508% |  | 18 | 18.0024 | 0.0024 | 0.01%
| 54613 | 20 | 20.0781 | 0.0781 | 0.3254% |  | 20 | 20.0115 | 0.0115 | 0.04792%
| 65535 | 24 | 24.1111 | 0.1111 | 0.4629% |  | 24 | 24.033 | 0.033 | 0.1375%
| Actual slope k | Actual intercept b | Transformation coefficient c | Transformation offset o value
Output 0-20mA test results | 0.0003061 | 0.0044508 | 0.99699 | -14.5404 |  |  | Calibration 0-20mA | Fitting the calibrated data results
measuring range | program parameter | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error |  | Theoretical output voltage/current (V/mA) | Actual output voltage/current (V/mA) | Error (V/mA) | full scale error
0-24mA（0-65535） | 0 | 0 | 0.0067 | 0.0067 | 0.02792% |  | 0 | 0 | 0 | 0%
| 6553.5 | 2 | 2.014 | 0.014 | 0.05833% |  | 2 | 2.0054 | 0.0054 | 0.0225%
| 13107 | 4 | 4.0197 | 0.0197 | 0.08208% |  | 4 | 4.0035 | 0.0035 | 0.01458%
| 26214 | 8 | 8.0276 | 0.0276 | 0.115% |  | 8 | 8 | 0 | 0%
| 32767.5 | 10 | 10.0259 | 0.0259 | 0.1079% |  | 10 | 10.0035 | 0.0035 | 0.01458%
| 39321 | 12 | 12.037 | 0.037 | 0.1542% |  | 12 | 11.997 | -0.003 | -0.0125%
| 52428 | 16 | 16.053 | 0.053 | 0.2208% |  | 16 | 16.0015 | 0.0015 | 0.00625%
| 58981.5 | 18 | 18.0651 | 0.0651 | 0.2713% |  | 18 | 18.0075 | 0.0075 | 0.03125%
| 65535 | 20 | 20.0813 | 0.0813 | 0.3387% |  | 20 | 20.0168 | 0.0168 | 0.07%